## calculations
Enter values in the gray cells, answers appear in yellow area
F0 | reading before acidification
Fa | reading after acidification |  |  |  | Calibration Constants:
Vol-filt | sample volume filtered (mL) |  |  |  | Kx | 0.4563 | (linear calibration factor)
Vol-ex | extraction volume (90% acetone, mL) |  |  |  | Fm | 1.9643 | (acidification coefficient) | 1.9643 | 2.2
blank | F0 | Fa | Vol-filt | Vol-ex | [chl a] (ug/L) | phaeo (chl equiv. Weights)
0.215 | 80 | 42 | 300 | 5 | 0.58868 | 0.035517
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215
0.215

## calibration
Calibration July 5, 2005
[chl a] (ug/L)
Stock solution: | 165 |  | blank: | 0.345
| dilution | fluorescence | fluorescence blank corrected | [chl a]
| 0 | 358 | 357.655 | 165
| 1 | 186 | 185.655 | 82.5
| 2 | 94 | 93.655 | 41.25
| 3 | 48.7 | 48.355 | 20.625
| 4 | 25 | 24.655 | 10.3125
| 5 | 12.9 | 12.555 | 5.1562
| blank | 0.345 | 0 | 0
|  | Kx |  | 0.4563
| dilution | fluorescence | fluorescence blank corrected | [chl a]
| 0 | 358 | 357.655 | 165
| 1 | 186 | 185.655 | 82.5
| 2 | 94 | 93.655 | 41.25
| 3 | 48.7 | 48.355 | 20.625
| 4 | 25 | 24.655 | 10.3125
| 5 | 12.9 | 12.555 | 5.1562
| 0 | 0.345 | 0 | 0
| 1ml + 5 ml | 71.4 | 71.055 | 27.5
| 1ml + 5 ml | 12.7 | 12.355 | 4.5833
| 2ml + 3ml | 146 | 145.655 | 66
|  | Kx |  | 0.4541
| F0 | Fa |  | F0/Fa
| 358 | 182 |  | 1.967
| 71.4 | 36.4 |  | 1.9615
| 146 | 76.6 |  | 1.906
| 106 | 55.8 |  | 1.8996
|  | Fm |  | 1.9643

## spectrum
Absorption Spectrum Chlorophyll a standard (165 ug/L)
| Absorption | [chl a] ug/L
peak | 0.014894 | 168.962 | absorption spectrum
664nm | 0.0149 | 169.0301 | wavelength abs
|  |  | 300 | 0.0058
|  |  | 301 | 0.0485
>extinction coefficient chl a at 664 nm = 88.15 au * L/g (using a 1 |  |  | 302 | 0.1748
>cm cuvette and 90% acetone) |  |  | 303 | -0.058
|  |  | 304 | 0.0164
|  |  | 305 | 0.2415
|  |  | 306 | -0.041
|  |  | 307 | 0.1485
|  |  | 308 | -0.08
|  |  | 309 | 0.1827
|  |  | 310 | 0.1944
|  |  | 311 | 0.111
|  |  | 312 | 0.132
|  |  | 313 | 0.0004
|  |  | 314 | 0.0966
|  |  | 315 | 0.1189
|  |  | 316 | 0.0461
|  |  | 317 | 0.0893
|  |  | 318 | 0.3092
|  |  | 319 | 0.6553
|  |  | 320 | 0.9075
|  |  | 321 | 1.0664
|  |  | 322 | 0.9533
|  |  | 323 | 0.8098
|  |  | 324 | 0.6637
|  |  | 325 | 0.5426
|  |  | 326 | 0.4378
|  |  | 327 | 0.3541
|  |  | 328 | 0.2812
|  |  | 329 | 0.2264
|  |  | 330 | 0.1828
|  |  | 331 | 0.1475
|  |  | 332 | 0.1196
|  |  | 333 | 0.0969
|  |  | 334 | 0.0792
|  |  | 335 | 0.065
|  |  | 336 | 0.054
|  |  | 337 | 0.0447
|  |  | 338 | 0.0377
|  |  | 339 | 0.0322
|  |  | 340 | 0.0281
|  |  | 341 | 0.0248
|  |  | 342 | 0.0222
|  |  | 343 | 0.0202
|  |  | 344 | 0.0188
|  |  | 345 | 0.0177
|  |  | 346 | 0.0166
|  |  | 347 | 0.0157
|  |  | 348 | 0.0151
|  |  | 349 | 0.0146
|  |  | 350 | 0.0143
|  |  | 351 | 0.014
|  |  | 352 | 0.0135
|  |  | 353 | 0.0133
|  |  | 354 | 0.013
|  |  | 355 | 0.0127
|  |  | 356 | 0.0125
|  |  | 357 | 0.0123
|  |  | 358 | 0.0121
|  |  | 359 | 0.0119
|  |  | 360 | 0.0117
|  |  | 361 | 0.0115
|  |  | 362 | 0.0115
|  |  | 363 | 0.0114
|  |  | 364 | 0.0113
|  |  | 365 | 0.0113
|  |  | 366 | 0.0112
|  |  | 367 | 0.0112
|  |  | 368 | 0.0111
|  |  | 369 | 0.0111
|  |  | 370 | 0.011
|  |  | 371 | 0.0112
|  |  | 372 | 0.0112
|  |  | 373 | 0.0112
|  |  | 374 | 0.0112
|  |  | 375 | 0.0112
|  |  | 376 | 0.0112
|  |  | 377 | 0.0112
|  |  | 378 | 0.0112
|  |  | 379 | 0.0111
|  |  | 380 | 0.0109
|  |  | 381 | 0.0109
|  |  | 382 | 0.0109
|  |  | 383 | 0.0108
|  |  | 384 | 0.0107
|  |  | 385 | 0.0106
|  |  | 386 | 0.0106
|  |  | 387 | 0.0103
|  |  | 388 | 0.0104
|  |  | 389 | 0.0103
|  |  | 390 | 0.0102
|  |  | 391 | 0.0102
|  |  | 392 | 0.0101
|  |  | 393 | 0.01
|  |  | 394 | 0.0103
|  |  | 395 | 0.0102
|  |  | 396 | 0.0104
|  |  | 397 | 0.0104
|  |  | 398 | 0.0106
|  |  | 399 | 0.0109
|  |  | 400 | 0.0113
|  |  | 401 | 0.0117
|  |  | 402 | 0.0119
|  |  | 403 | 0.0122
|  |  | 404 | 0.0127
|  |  | 405 | 0.013
|  |  | 406 | 0.0132
|  |  | 407 | 0.0134
|  |  | 408 | 0.0138
|  |  | 409 | 0.0138
|  |  | 410 | 0.0139
|  |  | 411 | 0.0139
|  |  | 412 | 0.014
|  |  | 413 | 0.0139
|  |  | 414 | 0.0138
|  |  | 415 | 0.0138
|  |  | 416 | 0.0138
|  |  | 417 | 0.0137
|  |  | 418 | 0.0136
|  |  | 419 | 0.0137
|  |  | 420 | 0.0137
|  |  | 421 | 0.0139
|  |  | 422 | 0.014
|  |  | 423 | 0.0145
|  |  | 424 | 0.015
|  |  | 425 | 0.0154
|  |  | 426 | 0.016
|  |  | 427 | 0.0165
|  |  | 428 | 0.0171
|  |  | 429 | 0.0175
|  |  | 430 | 0.0178
|  |  | 431 | 0.0179
|  |  | 432 | 0.018
|  |  | 433 | 0.0176
|  |  | 434 | 0.0172
|  |  | 435 | 0.0164
|  |  | 436 | 0.0155
|  |  | 437 | 0.0142
|  |  | 438 | 0.0129
|  |  | 439 | 0.0116
|  |  | 440 | 0.0103
|  |  | 441 | 0.009
|  |  | 442 | 0.0081
|  |  | 443 | 0.0069
|  |  | 444 | 0.0058
|  |  | 445 | 0.0049
|  |  | 446 | 0.0043
|  |  | 447 | 0.0037
|  |  | 448 | 0.0033
|  |  | 449 | 0.0029
|  |  | 450 | 0.0025
|  |  | 451 | 0.0022
|  |  | 452 | 0.0022
|  |  | 453 | 0.0018
|  |  | 454 | 0.0018
|  |  | 455 | 0.0016
|  |  | 456 | 0.0015
|  |  | 457 | 0.0015
|  |  | 458 | 0.0014
|  |  | 459 | 0.0013
|  |  | 460 | 0.0012
|  |  | 461 | 0.0013
|  |  | 462 | 0.0011
|  |  | 463 | 0.0012
|  |  | 464 | 0.0011
|  |  | 465 | 0.001
|  |  | 466 | 0.0011
|  |  | 467 | 0.0011
|  |  | 468 | 0.0011
|  |  | 469 | 0.0011
|  |  | 470 | 0.001
|  |  | 471 | 0.0012
|  |  | 472 | 0.0011
|  |  | 473 | 0.0011
|  |  | 474 | 0.0011
|  |  | 475 | 0.0011
|  |  | 476 | 0.0011
|  |  | 477 | 0.0011
|  |  | 478 | 0.0009
|  |  | 479 | 0.0011
|  |  | 480 | 0.0009
|  |  | 481 | 0.0011
|  |  | 482 | 0.001
|  |  | 483 | 0.001
|  |  | 484 | 0.0009
|  |  | 485 | 0.001
|  |  | 486 | 0.001
|  |  | 487 | 0.001
|  |  | 488 | 0.001
|  |  | 489 | 0.001
|  |  | 490 | 0.001
|  |  | 491 | 0.0011
|  |  | 492 | 0.0009
|  |  | 493 | 0.001
|  |  | 494 | 0.001
|  |  | 495 | 0.0011
|  |  | 496 | 0.001
|  |  | 497 | 0.001
|  |  | 498 | 0.001
|  |  | 499 | 0.0011
|  |  | 500 | 0.001
|  |  | 501 | 0.0011
|  |  | 502 | 0.001
|  |  | 503 | 0.001
|  |  | 504 | 0.001
|  |  | 505 | 0.0009
|  |  | 506 | 0.0009
|  |  | 507 | 0.0012
|  |  | 508 | 0.0009
|  |  | 509 | 0.0009
|  |  | 510 | 0.001
|  |  | 511 | 0.0009
|  |  | 512 | 0.001
|  |  | 513 | 0.0012
|  |  | 514 | 0.0011
|  |  | 515 | 0.0011
|  |  | 516 | 0.0011
|  |  | 517 | 0.001
|  |  | 518 | 0.0012
|  |  | 519 | 0.0011
|  |  | 520 | 0.0011
|  |  | 521 | 0.0011
|  |  | 522 | 0.0011
|  |  | 523 | 0.0012
|  |  | 524 | 0.0012
|  |  | 525 | 0.0012
|  |  | 526 | 0.0012
|  |  | 527 | 0.0011
|  |  | 528 | 0.0012
|  |  | 529 | 0.0012
|  |  | 530 | 0.0012
|  |  | 531 | 0.0012
|  |  | 532 | 0.0013
|  |  | 533 | 0.0012
|  |  | 534 | 0.0012
|  |  | 535 | 0.0013
|  |  | 536 | 0.0013
|  |  | 537 | 0.0013
|  |  | 538 | 0.0013
|  |  | 539 | 0.0012
|  |  | 540 | 0.0012
|  |  | 541 | 0.0011
|  |  | 542 | 0.0013
|  |  | 543 | 0.0012
|  |  | 544 | 0.0012
|  |  | 545 | 0.0012
|  |  | 546 | 0.0012
|  |  | 547 | 0.0011
|  |  | 548 | 0.0011
|  |  | 549 | 0.0012
|  |  | 550 | 0.0013
|  |  | 551 | 0.001
|  |  | 552 | 0.0011
|  |  | 553 | 0.0011
|  |  | 554 | 0.0012
|  |  | 555 | 0.0013
|  |  | 556 | 0.0013
|  |  | 557 | 0.0013
|  |  | 558 | 0.0013
|  |  | 559 | 0.0013
|  |  | 560 | 0.0013
|  |  | 561 | 0.0014
|  |  | 562 | 0.0015
|  |  | 563 | 0.0014
|  |  | 564 | 0.0016
|  |  | 565 | 0.0015
|  |  | 566 | 0.0016
|  |  | 567 | 0.0017
|  |  | 568 | 0.0017
|  |  | 569 | 0.0018
|  |  | 570 | 0.0019
|  |  | 571 | 0.0017
|  |  | 572 | 0.002
|  |  | 573 | 0.0018
|  |  | 574 | 0.002
|  |  | 575 | 0.0019
|  |  | 576 | 0.0019
|  |  | 577 | 0.0019
|  |  | 578 | 0.002
|  |  | 579 | 0.0021
|  |  | 580 | 0.0022
|  |  | 581 | 0.0021
|  |  | 582 | 0.0022
|  |  | 583 | 0.0021
|  |  | 584 | 0.0022
|  |  | 585 | 0.0021
|  |  | 586 | 0.0021
|  |  | 587 | 0.0022
|  |  | 588 | 0.002
|  |  | 589 | 0.0019
|  |  | 590 | 0.0019
|  |  | 591 | 0.0019
|  |  | 592 | 0.0019
|  |  | 593 | 0.0019
|  |  | 594 | 0.0019
|  |  | 595 | 0.002
|  |  | 596 | 0.0019
|  |  | 597 | 0.002
|  |  | 598 | 0.0021
|  |  | 599 | 0.0022
|  |  | 600 | 0.0022
|  |  | 601 | 0.0023
|  |  | 602 | 0.0023
|  |  | 603 | 0.0024
|  |  | 604 | 0.0025
|  |  | 605 | 0.0026
|  |  | 606 | 0.0028
|  |  | 607 | 0.0027
|  |  | 608 | 0.0028
|  |  | 609 | 0.003
|  |  | 610 | 0.0031
|  |  | 611 | 0.0031
|  |  | 612 | 0.0031
|  |  | 613 | 0.0031
|  |  | 614 | 0.003
|  |  | 615 | 0.0032
|  |  | 616 | 0.0032
|  |  | 617 | 0.0032
|  |  | 618 | 0.0033
|  |  | 619 | 0.0032
|  |  | 620 | 0.0032
|  |  | 621 | 0.0032
|  |  | 622 | 0.0032
|  |  | 623 | 0.0031
|  |  | 624 | 0.003
|  |  | 625 | 0.003
|  |  | 626 | 0.0029
|  |  | 627 | 0.0028
|  |  | 628 | 0.0027
|  |  | 629 | 0.0026
|  |  | 630 | 0.0026
|  |  | 631 | 0.0025
|  |  | 632 | 0.0025
|  |  | 633 | 0.0025
|  |  | 634 | 0.0024
|  |  | 635 | 0.0023
|  |  | 636 | 0.0022
|  |  | 637 | 0.0023
|  |  | 638 | 0.0023
|  |  | 639 | 0.0024
|  |  | 640 | 0.0025
|  |  | 641 | 0.0025
|  |  | 642 | 0.0026
|  |  | 643 | 0.0028
|  |  | 644 | 0.0029
|  |  | 645 | 0.0033
|  |  | 646 | 0.0034
|  |  | 647 | 0.0039
|  |  | 648 | 0.0043
|  |  | 649 | 0.0047
|  |  | 650 | 0.0053
|  |  | 651 | 0.006
|  |  | 652 | 0.0068
|  |  | 653 | 0.0075
|  |  | 654 | 0.0084
|  |  | 655 | 0.0093
|  |  | 656 | 0.0101
|  |  | 657 | 0.0111
|  |  | 658 | 0.012
|  |  | 659 | 0.0128
|  |  | 660 | 0.0137
|  |  | 661 | 0.0142
|  |  | 662 | 0.0147
|  |  | 663 | 0.0149
|  |  | 664 | 0.0149
|  |  | 665 | 0.0147
|  |  | 666 | 0.0141
|  |  | 667 | 0.0135
|  |  | 668 | 0.0127
|  |  | 669 | 0.0118
|  |  | 670 | 0.0106
|  |  | 671 | 0.0096
|  |  | 672 | 0.0086
|  |  | 673 | 0.0072
|  |  | 674 | 0.0065
|  |  | 675 | 0.0057
|  |  | 676 | 0.0047
|  |  | 677 | 0.0041
|  |  | 678 | 0.0036
|  |  | 679 | 0.003
|  |  | 680 | 0.0025
|  |  | 681 | 0.0021
|  |  | 682 | 0.0018
|  |  | 683 | 0.0016
|  |  | 684 | 0.0014
|  |  | 685 | 0.0011
|  |  | 686 | 0.001
|  |  | 687 | 0.0009
|  |  | 688 | 0.0007
|  |  | 689 | 0.0007
|  |  | 690 | 0.0007
|  |  | 691 | 0.0006
|  |  | 692 | 0.0006
|  |  | 693 | 0.0004
|  |  | 694 | 0.0006
|  |  | 695 | 0.0005
|  |  | 696 | 0.0005
|  |  | 697 | 0.0005
|  |  | 698 | 0.0004
|  |  | 699 | 0.0005
|  |  | 700 | 0.0004
|  |  | 701 | 0.0003
|  |  | 702 | 0.0004
|  |  | 703 | 0.0003
|  |  | 704 | 0.0004
|  |  | 705 | 0.0004
|  |  | 706 | 0.0003
|  |  | 707 | 0.0004
|  |  | 708 | 0.0003
|  |  | 709 | 0.0003
|  |  | 710 | 0.0003
|  |  | 711 | 0.0003
|  |  | 712 | 0.0003
|  |  | 713 | 0.0001
|  |  | 714 | 0.0003
|  |  | 715 | 0.0002
|  |  | 716 | 0.0002
|  |  | 717 | 0.0001
|  |  | 718 | 0.0001
|  |  | 719 | 0.0002
|  |  | 720 | 0
|  |  | 721 | 0.0002
|  |  | 722 | 0
|  |  | 723 | 0.0001
|  |  | 724 | 0.0001
|  |  | 725 | 0
|  |  | 726 | 0.0002
|  |  | 727 | 0
|  |  | 728 | 0.0002
|  |  | 729 | 0
|  |  | 730 | 0
|  |  | 731 | 0
|  |  | 732 | 0
|  |  | 733 | 0
|  |  | 734 | 0
|  |  | 735 | 0
|  |  | 736 | 0
|  |  | 737 | 0
|  |  | 738 | 0
|  |  | 739 | 0
|  |  | 740 | 0
|  |  | 741 | 0
|  |  | 742 | 0
|  |  | 743 | 0
|  |  | 744 | 0
|  |  | 745 | 0
|  |  | 746 | 0
|  |  | 747 | 0
|  |  | 748 | 0
|  |  | 749 | 0
|  |  | 750 | 0
|  |  | 751 | 0
|  |  | 752 | -0.001
|  |  | 753 | 0
|  |  | 754 | -0.001
|  |  | 755 | 0
|  |  | 756 | -0.001
|  |  | 757 | 0
|  |  | 758 | -0.001
|  |  | 759 | 0
|  |  | 760 | -0.001
|  |  | 761 | -0.001
|  |  | 762 | 0
|  |  | 763 | -0.001
|  |  | 764 | 0
|  |  | 765 | -0.001
|  |  | 766 | 0
|  |  | 767 | -0.001
|  |  | 768 | 0
|  |  | 769 | -0.001
|  |  | 770 | 0
|  |  | 771 | 0
|  |  | 772 | 0
|  |  | 773 | 0
|  |  | 774 | 0
|  |  | 775 | 0
|  |  | 776 | 0
|  |  | 777 | 0
|  |  | 778 | -0.001
|  |  | 779 | 0
|  |  | 780 | 0
|  |  | 781 | 0
|  |  | 782 | 0
|  |  | 783 | 0
|  |  | 784 | 0
|  |  | 785 | 0
|  |  | 786 | 0
|  |  | 787 | 0
|  |  | 788 | 0
|  |  | 789 | 0
|  |  | 790 | 0
|  |  | 791 | 0
|  |  | 792 | 0
|  |  | 793 | 0
|  |  | 794 | 0
|  |  | 795 | 0
|  |  | 796 | 0
|  |  | 797 | 0
|  |  | 798 | 0
|  |  | 799 | 0
|  |  | 800 | 0

## jgofs_protocol
JGOFS Protocol:
see also: | http://ocean.stanford.edu/cal/JGOFS_chla.pdf
JGOFS Protocols—June 1994 119
Chapter 14. Measurement of Chlorophyll a and Phaeopigments by
Fluorometric Analysis
1.0 Scope and field of application
Chlorophyll a measurements have historically provided a useful estimate of algal biomass
and its spatial and temporal variability. The fluorometric method is extensively used for
the quantitative analysis of chlorophyll a and phaeopigments. However, errors can be
introduced into the results when chlorophylls b and/or chlorophylls c are present.
Chlorophyll b is the main source of error in this method. While generally not abundant in
surface waters, chlorophyll b can be as high as 0.5 times the chlorophyll a concentration in
the deep chlorophyll maximum, causing slight underestimations of the chlorophyll a
concentration, and drastic overestimations of the phaeopigment concentrations. Divinyl-chlorophyll
a also interferes and is taken as chlorophyll a by this method. The procedure
described here is appropriate for all levels of chlorophyll a concentration in the marine
environment. Filtration volumes should be modified for the different environments.
Scientists who employ this or other methods to measure pigments should make themselves
aware of the current and historical issues that surround these techniques and make
appropriate decisions about specific methodologies for their application based on the
scientific requirements and constraints of their individual programs.
2.0 Definition
The concentrations of chlorophyll a and phaeopigments in seawater are given as
mg kg -1 .
3.0 Principle of Analysis
Algal pigments, particularly chlorophyll a, fluoresce in the red wavelengths after
extraction in acetone when they are excited by blue wavelengths of light. The fluorometer
excites the extracted sample with a broadband blue light and the resulting fluorescence in
the red is detected by a photomultiplier. The significant fluorescence by phaeopigments is
corrected for by acidifying the sample which converts all of the chlorophyll a to
phaeopigments. By applying a measured conversion for the relative strength of
chlorophyll and phaeopigment fluorescence, the two values can be used to calculate both
the chlorophyll a and phaeopigment concentrations.
4.0 Apparatus
4.1 Filtration system and Whatman GF/F filters
4.2 Liquid nitrogen and freezer for storage and extraction
4.3 Glass centrifuge tubes for extraction, 15 ml
4.4 Turner fluorometer, fitted with a red sensitive photomultiplier, a blue lamp, 5-60 blue
filter and 2-64 red filter.
5.0 Reagents
5.1 100% acetone
5.2 90% acetone
5.3 1.2M HCl (100 ml HCl in 900 ml de-ionized water)
6.0 Sample Collection and Storage
Water samples are collected from niskins into clean polyethylene bottles with Tygon Ò
tubing. Samples are immediately filtered through 47 mm GF/F filters using polycarbonate
in-line filters (Gelman) and a vacuum of less than 100 mm Hg. Filters are folded in half
twice and wrapped in aluminum foil, labeled, and stored in liquid nitrogen (to avoid
formation of degradation products) until shore analysis. Alternatively, filters can be placed
immediately in acetone for pigment extraction if analysis is to be carried out onboard ship.
In oligotrophic waters, for this measurement coupled with HPLC determined pigments, 4
liters are filtered. For fluorometric analysis alone, a smaller volume (0.5 -1.0 l) may be
sufficient. In coastal regions, a volume of 0.1-0.5 l may be adequate. In this case, use of 25
mm GF/F filters may be appropriate.
7.0 Procedure
7.1 After removal from liquid nitrogen or freezer), the pigments are extracted by placing
the filters in 5.0 ml 100% acetone. For 47 mm GF/F filters, 0.8 ml of water is
retained adjusting the final extraction solution to 86% acetone and the final extrac-tion
volume to 5.8 ml. The samples are covered with Parafilm to reduce evaporation,
sonicated (0°C, subdued light) and allowed to extract for 4 hours in the dark at
-20°C. Following extraction, samples are vortexed, filters are pressed to the bottom
of the tube with a stainless steel spatula and spun down in a centrifuge for 5 minutes
to remove cellular debris. For fluorometric analysis (not HPLC), decantation can
replace centrifuging.
7.1.1 The addition of 5.0 ml acetone for pigment extraction is necessary to com-pletely
submerge 47 mm GF/F filters in 15 ml centrifuge tubes. This volume
may be altered depending on the size of the filter and volume of the extrac-tion
tube.
7.2 The fluorometer is allowed to warm up and stabilize for 30 minutes prior to use.
7.3 The fluorometer is zeroed with 90% acetone.
7.4 1.0 ml of pigment extract is mixed with 4.0 ml 90% acetone in a cuvette and read on
the appropriate door to give a reading between 30 and 100. The sample is then acidi-fied
with 2 drops of 1.2 M HCl. Further dilutions may be necessary for higher chlo-rophyll
a concentrations.
7.5 Standardization
7.5.1 For laboratory use, the fluorometer is calibrated every 6 months with a com-mercially
available chlorophyll a standard (Anacystis nidulans, Sigma
Chemical Company). If the fluorometer is taken to sea, it is recommended
that the fluorometer be calibrated before and after each cruise.
7.5.2 The standard is dissolved in 90% acetone for at least 2 hours and it’s concen-tration
(mg l -1 ) is calculated spectrophotometrically as follows:
where:
A max = absorption maximum (664 nm)
A 750 nm = absorbance at 750 nm to correct for light scattering
E = extinction coefficient for chl a in 90% acetone at 664 nm
(87.67 L g -1 cm -1 )
l = cuvette path length (cm)
7.5.3 From the standard, a minimum of five dilutions are prepared for each door.
Fluorometer readings are taken before and after acidification with 2 drops
1.2 M HCl.
7.5.4 Linear calibration factor (K x ) are calculated for each door (x) as the slope of
the unacidified fluorometric reading vs. chlorophyll a concentration calcu-lated
spectrophotometrically.
chla A max A 750nm – ( )
E l ´ --------------------------------------- 1000mg
1gram ------------------- - ´ =
122 JGOFS Protocols—June 1994
7.5.5 The acidification coefficient (F m ) is calculated by averaging the ratio of the
unacidified and acidified readings (F o /F a ) of pure chlorophyll a.
7.5.6 Samples are read using a door setting that produces a dial reading between 30
and 100. The fluorometer is zeroed with 90% acetone each time the door set-ting
is changed.
8.0 Calculation and expression of results
The concentrations of chlorophyll a and phaeopigments in the sample are calculated using
the following equations:
where:
F m = acidification coefficient (F o /F a ) for pure Chl a (usually 2.2).
F o = reading before acidification
F a = reading after acidification
K x = door factor from calibration calculations
vol ex = extraction volume
vol filt = sample volume
9.0 References
Herbland, A., A. Le Bouteiller, and P. Raimbault. (1985). Size structure of phytoplankton
biomass in the equatorial Atlantic Ocean. Deep-Sea Res., 32: 819-836.
Holm-Hansen, O., and B. Riemann. (1978). Chlorophyll a determination: improvements
in methodology. Oikos, 30: 438-447.
Chl (mg/l) F m
F m
------------- è –1ø æ ö F o a –F ( ) ´ K x
vol ex
vol filt
------------ - è ø æ ö ´ ´ =
Phaeo (chl equiv. weights) F m
F m
------------- è –1ø æ ö F m F a · ( o )–F [ ]K x ex –vol ´ =

## Sheet2
| 1 ml concentration | in 1 ml | total 6 ml
1 ml in 5 ml acetone | 165 | 0.165 | 0.0275 | 27.5
| ug/l | ug | ug/ml
| 27.5 | 0.0275 | 0.0045833 | 4.5833
2 ml
2 ml + 3 ml | 165 | 0.33 | 0.066 | 66